CALIBRATION CERTIFICATE

Date of calibration:	19.02.2026
Type:	X-ray radiation dosimeter AT1103M
S/N:	000
Measurement limits:	X: 50 nSv/h – 100 µSv/h
Measurement error:	±15%

Operating conditions:
- Air temperature	+19.0 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	72.0 %
- Gamma radiation background	98.0 nSv/h

Calibration means:
- the standard dosimetry facility AT110, № 013 reg. №40425-09, uncertainty ±5% (the Calibration Certificate №С-В/26-11-2024/409219101 on 26 November 2024 issued by The D.I. Mendeleev All-Russian Institute for Metrology (VNIIM), St. Petersburg, Russia);

Calibration data

X-ray radiation dosimeter AT1103M №000
Conventionally true value of the dose rate from 241Am radionuclide, {{H}} rsub {0 i} rsup {'} \(0.07\)	Calibration data / {bar {H rsub {i} rsup {"'"}}} (10)
0.07 µSv/h	0.0703 µSv/h
0.7 µSv/h	0.713 µSv/h
7.0 µSv/h	7.023 µSv/h
70.0 µSv/h	69.73 µSv/h

The radiation dosimeter complies with the measurement specifications in the User’s Manual.

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)