CALIBRATION CERTIFICATE
AT2327 Alarm Dosimeter	Date of calibration: $date.

SN: $serial

Type: BDKN-02

Neutron flux density measurement limits: BDKN-01: 0.1 – 104 s-1cm-2 ;

Measurement error: BDKN-02: ±20%

Operating conditions:
- Air temperature	+$temp °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	$vlag %
- Gamma radiation background	$bground nSv/h

Calibration means:
$eng_facility_140;

Calibration data

BDKN-02 (n) s/n: $bd_num1
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$1-61 / $1-71 / $1-81	$1-62 / $1-72 / $1-82	$1-63 / $1-73 / $1-83	$1-64 / $1-74 / $1-84
Average readings, i, s-1cm-2	$1-00	$1-01	$1-02	$1-03	$1-04
Result of flux density measurement at control point / npi, s-1cm-2		$1-n21	$1-n22	$1-n23	$1-n24
Intrinsic relative error of measurement, θпрi, %		$1-n31	$1-n32	$1-n33	$1-n34
Confidence limit of the intrinsic relative error Δi, %		$1-n41	$1-n42	$1-n43	$1-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num2
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$2-61 / $2-71 / $2-81	$2-62 / $2-72 / $2-82	$2-63 / $2-73 / $2-83	$2-64 / $2-74 / $2-84
Average readings, i, s-1cm-2	$2-00	$2-01	$2-02	$2-03	$2-04
Result of flux density measurement at control point / npi, s-1cm-2		$2-n21	$2-n22	$2-n23	$2-n24
Intrinsic relative error of measurement, θпрi, %		$2-n31	$2-n32	$2-n33	$2-n34
Confidence limit of the intrinsic relative error Δi, %		$2-n41	$2-n42	$2-n43	$2-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num3
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$3-61 / $3-71 / $3-81	$3-62 / $3-72 / $3-82	$3-63 / $3-73 / $3-83	$3-64 / $3-74 / $3-84
Average readings, i, s-1cm-2	$3-00	$3-01	$3-02	$3-03	$3-04
Result of flux density measurement at control point / npi, s-1cm-2		$3-n21	$3-n22	$3-n23	$3-n24
Intrinsic relative error of measurement, θпрi, %		$3-n31	$3-n32	$3-n33	$3-n34
Confidence limit of the intrinsic relative error Δi, %		$3-n41	$3-n42	$3-n43	$3-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num4
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$4-61 / $4-71 / $4-81	$4-62 / $4-72 / $4-82	$4-63 / $4-73 / $4-83	$4-64 / $4-74 / $4-84
Average readings, i, s-1cm-2	$4-00	$4-01	$4-02	$4-03	$4-04
Result of flux density measurement at control point / npi, s-1cm-2		$4-n21	$4-n22	$4-n23	$4-n24
Intrinsic relative error of measurement, θпрi, %		$4-n31	$4-n32	$4-n33	$4-n34
Confidence limit of the intrinsic relative error Δi, %		$4-n41	$4-n42	$4-n43	$4-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num5
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$5-61 / $5-71 / $5-81	$5-62 / $5-72 / $5-82	$5-63 / $5-73 / $5-83	$5-64 / $5-74 / $5-84
Average readings, i, s-1cm-2	$5-00	$5-01	$5-02	$5-03	$5-04
Result of flux density measurement at control point / npi, s-1cm-2		$5-n21	$5-n22	$5-n23	$5-n24
Intrinsic relative error of measurement, θпрi, %		$5-n31	$5-n32	$5-n33	$5-n34
Confidence limit of the intrinsic relative error Δi, %		$5-n41	$5-n42	$5-n43	$5-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num6
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$6-61 / $6-71 / $6-81	$6-62 / $6-72 / $6-82	$6-63 / $6-73 / $6-83	$6-64 / $6-74 / $6-84
Average readings, i, s-1cm-2	$6-00	$6-01	$6-02	$6-03	$6-04
Result of flux density measurement at control point / npi, s-1cm-2		$6-n21	$6-n22	$6-n23	$6-n24
Intrinsic relative error of measurement, θпрi, %		$6-n31	$6-n32	$6-n33	$6-n34
Confidence limit of the intrinsic relative error Δi, %		$6-n41	$6-n42	$6-n43	$6-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num7
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$7-61 / $7-71 / $7-81	$7-62 / $7-72 / $7-82	$7-63 / $7-73 / $7-83	$7-64 / $7-74 / $7-84
Average readings, i, s-1cm-2	$7-00	$7-01	$7-02	$7-03	$7-04
Result of flux density measurement at control point / npi, s-1cm-2		$7-n21	$7-n22	$7-n23	$7-n24
Intrinsic relative error of measurement, θпрi, %		$7-n31	$7-n32	$7-n33	$7-n34
Confidence limit of the intrinsic relative error Δi, %		$7-n41	$7-n42	$7-n43	$7-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num8
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$8-61 / $8-71 / $8-81	$8-62 / $8-72 / $8-82	$8-63 / $8-73 / $8-83	$8-64 / $8-74 / $8-84
Average readings, i, s-1cm-2	$8-00	$8-01	$8-02	$8-03	$8-04
Result of flux density measurement at control point / npi, s-1cm-2		$8-n21	$8-n22	$8-n23	$8-n24
Intrinsic relative error of measurement, θпрi, %		$8-n31	$8-n32	$8-n33	$8-n34
Confidence limit of the intrinsic relative error Δi, %		$8-n41	$8-n42	$8-n43	$8-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num9
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$9-61 / $9-71 / $9-81	$9-62 / $9-72 / $9-82	$9-63 / $9-73 / $9-83	$9-64 / $9-74 / $9-84
Average readings, i, s-1cm-2	$9-00	$9-01	$9-02	$9-03	$9-04
Result of flux density measurement at control point / npi, s-1cm-2		$9-n21	$9-n22	$9-n23	$9-n24
Intrinsic relative error of measurement, θпрi, %		$9-n31	$9-n32	$9-n33	$9-n34
Confidence limit of the intrinsic relative error Δi, %		$9-n41	$9-n42	$9-n43	$9-n44
Limite of intrinsic relative error , % not above		±20

BDKN-02 (n) s/n: $bd_num10
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$10-61 / $10-71 / $10-81	$10-62 / $10-72 / $10-82	$10-63 / $10-73 / $10-83	$10-64 / $10-74 / $10-84
Average readings, i, s-1cm-2	$10-00	$10-01	$10-02	$10-03	$10-04
Result of flux density measurement at control point / npi, s-1cm-2		$10-n21	$10-n22	$10-n23	$10-n24
Intrinsic relative error of measurement, θпрi, %		$10-n31	$10-n32	$10-n33	$10-n34
Confidence limit of the intrinsic relative error Δi, %		$10-n41	$10-n42	$10-n43	$10-n44
Limite of intrinsic relative error , % not above		±20

Calibrated by:	$calibratedBy
		(signature)
Technical control:	$controlledBy
		(signature)