CALIBRATION CERTIFICATE
Type: Radiation monitor $type	Date of calibration: $date.

S/N: $serial

Measurement limits of ambient x-ray gamma radiation:

АТ1125: γ 30 nSv/h – 300 µSv/h;

АТ1125А: γ 30 nSv/h – 100 mSv/h;

BDPS-02: γ 0.1 µSv/h – 30.0 mSv/h; / Flux density measurement limits of α particles: / BDPS-02: α 2.4 – 106 min-1⋅cm-2; / Flux density measurement limits of β particles: / BDPS-02: β 6.0 – 106 min-1⋅cm-2;

SA of 137Cs count rate: / from $ua_min s-1 to 105 s-1

03: ± 20 %БД
Measurement error:
АТ1125: ± 15 %;

АТ1125А: ± 15 %;

BDPS-02: α 2,4 – 30 min-1⋅ sm- 2 – ± 30 % / α 30 – 106 min-1⋅ sm- 2 – ± 20 % / β 6 – 106 min-1⋅ sm- 2 – ± 20 % / γ 0.10 µSv/h – 30 mSv/h – ± 20 %

SA of 137Cs count rate: ± 20 %

Operating conditions:
- Air temperature	+$temp °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	$vlag %
- Gamma radiation background	$bground nSv/h

Calibration means:
- $eng_facility_110$;

- $eng_facility_130$;

- $eng_facility_alpha$;

- $eng_facility_beta$;

Calibration data / Radiation monitor $type s/n $serial
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,24 µSv/h	$1-30	$1-40	$1-50	$1-60	$1-70	$1-80	$1-00	$1-10	$1-20	±15
0,7 µSv/h	$1-31	$1-41	$1-51	$1-61	$1-71	$1-81	$1-01	$1-11	$1-21
7 µSv/h	$1-32	$1-42	$1-52	$1-62	$1-72	$1-82	$1-02	$1-12	$1-22
70 µSv/h	$1-33	$1-43	$1-53	$1-63	$1-73	$1-83	$1-03	$1-13	$1-23
240 µSv/h	$1-34	$1-44	$1-54	$1-64	$1-74	$1-84	$1-04	$1-14	$1-24

Calibration data / Radiation monitor 1125A s/n $serial
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,24 µSv/h	$2-30	$2-40	$2-50	$2-60	$2-70	$2-80	$2-00	$2-10	$2-20	±15
0,7 µSv/h	$2-31	$2-41	$2-51	$2-61	$2-71	$2-81	$2-01	$2-11	$2-21
7 µSv/h	$2-32	$2-42	$2-52	$2-62	$2-72	$2-82	$2-02	$2-12	$2-22
70 µSv/h	$2-33	$2-43	$2-53	$2-63	$2-73	$2-83	$2-03	$2-13	$2-23
240 µSv/h	$2-34	$2-44	$2-54	$2-64	$2-74	$2-84	$2-04	$2-14	$2-24
0,7 mSv/h	$2-35	$2-45	$2-55	$2-65	$2-75	$2-85	$2-05	$2-15	$2-25
7 mSv/h	$2-36	$2-46	$2-56	$2-66	$2-76	$2-86	$2-06	$2-16	$2-26
70 mSv/h	$2-37	$2-47	$2-57	$2-67	$2-77	$2-87	$2-07	$2-17	$2-27

SA of 137Cs
Radia-tion source number	SA in reference point, / Aoi, Bq/kg	Dose rate measurement at check point				Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Measured value Ḣ*i(10)			Average value, Aoi
		Ao1	Ao2	Ao3
$6-A0	$6-90	$6-60	$6-70	$6-80	$6-00	$6-10	$6-20	±20
$6-A1	$6-91	$6-61	$6-71	$6-81	$6-01	$6-11	$6-21
$6-A2	$6-92	$6-62	$6-72	$6-82	$6-02	$6-12	$6-22
$6-A3	$6-93	$6-63	$6-73	$6-83	$6-03	$6-13	$6-23

SA of 137Cs
Radia-tion source number	SA in reference point, / Aoi, Bq/kg	Dose rate measurement at check point				Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Measured value Ḣ*i(10)			Average value, Aoi
		Ao1	Ao2	Ao3
$6-A1	$6-91	$6-61	$6-71	$6-81	$6-01	$6-11	$6-21	±20
$6-A2	$6-92	$6-62	$6-72	$6-82	$6-02	$6-12	$6-22
$6-A3	$6-93	$6-63	$6-73	$6-83	$6-03	$6-13	$6-23

BDPS-02 (a) s/n $serial
Flux density in test point φ0i, / min-1·cm-2	Radia-tion source number	Measured radiation flux density in test point, φi				Relative alpha radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Measured value φi			Average value, φ
		φ1	φ2	φ3
$3-90	1179	$3-60	$3-70	$3-80	$3-00	$3-10	$3-20	±20
$3-91	1180	$3-61	$3-71	$3-81	$3-01	$3-11	$3-21
$3-92	1181	$3-62	$3-72	$3-82	$3-02	$3-12	$3-22
$3-93	1182	$3-63	$3-73	$3-83	$3-03	$3-13	$3-23
$3-94	1183	$3-64	$3-74	$3-84	$3-04	$3-14	$3-24

BDPS-02 (b) s/n $serial
Flux density in test point φ0i, / min-1·cm-2	Radia-tion source number	Measured radiation flux density in test point, φi				Relative beta radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Measured value φi			Average value, φ
		φ1	φ2	φ3
$4-90	0660	$4-60	$4-70	$4-80	$4-00	$4-10	$4-20	±20
$4-91	6905	$4-61	$4-71	$4-81	$4-01	$4-11	$4-21
$4-92	0661	$4-62	$4-72	$4-82	$4-02	$4-12	$4-22
$4-93	0662	$4-63	$4-73	$4-83	$4-03	$4-13	$4-23
$4-94	15014	$4-64	$4-74	$4-84	$4-04	$4-14	$4-24

BDPS-02 (γ) s/n $serial
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,7 µSv/h	$5-30	$5-40	$5-50	$5-60	$5-70	$5-80	$5-00	$5-10	$5-20	±20
7 µSv/h	$5-31	$5-41	$5-51	$5-61	$5-71	$5-81	$5-01	$5-11	$5-21
70 µSv/h	$5-32	$5-42	$5-52	$5-62	$5-72	$5-82	$5-02	$5-12	$5-22
0.7 mSv/h	$5-33	$5-43	$5-53	$5-63	$5-73	$5-83	$5-03	$5-13	$5-23
7,0 mSv/h	$5-34	$5-44	$5-54	$5-64	$5-74	$5-84	$5-04	$5-14	$5-24
20 mSv/h	$5-35	$5-45	$5-55	$5-65	$5-75	$5-85	$5-05	$5-15	$5-25

Calibrated by:	$calibratedBy
		(signature)
Technical control:	$controlledBy
		(signature)