CALIBRATION CERTIFICATE

Type: Personal dosimeter AT2503B	Date of calibration: $date

S/n: $allSerial

Measurement range for:
- individual dose equivalent:	0.1 µSv – 10 Sv
- individual dose equivalent rate:	0.1 µSv/h – 1 Sv/h
Measurement error:
- dose response variation limits	±(15+UrelD)%
- dose rate response variation limits	±(15+UrelD)%
where UrelD is the relative expanded uncertainty of the conventional value of a dose as a percentage
- dose coefficient of variation (H<11H0)	±(16-H/H0)%
- dose coefficient of variation (H⩾11H0)	±5%
- dose rate coefficient of variation (Ḣ<10 µSv/h)	±20%
- dose rate coefficient of variation (10µSv/h≤Ḣ<60µSv/h)	±(21-Ḣ/10µSv/h)%
- dose rate coefficient of variation (Ḣ⩾60µSv/h)	±15%
where H0 is the lower limit of dose measurement

Operating conditions:
- air temperature	+$temp °C
- relative humidity	$vlag %
- gamma radiation background	$bground nSv/h

Calibration means:
- $eng_facility_110;
- $eng_facility_130;
Calibration data:
AT2503B (γ) s/n: $bd_num1$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement / at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, / %
Background	—	$1-d0		$1-av-bg	—	—	—	—
0.80 µSv/h	$1-n0	$1-d0		$1-av0	$1-o0	±20	$1-q0	±20
4.0 mSv/h	$1-n1	$1-d1		$1-av1	$1-o1	±19	$1-q1	±15
80.0 mSv/h	$1-n2	$1-d2		$1-av2	$1-o2	±19	$1-q2	±15
400 mSv/h	$1-n3	$1-d3		$1-av3	$1-o3	±19	$1-q3	±15
800 mSv/h	$1-n4	$1-d4		$1-av4	$1-o4	±19	$1-q4	±15

AT2503B (γ) s/n: $bd_num1$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$1-dn_1 / $1-dd_1	$1-d_dev_1	$1-dose_1	$1-d_res_1	$1-o_pog_1	±20	$1-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$1-dn_2 / $1-dd_2	$1-d_dev_2	$1-dose_2	$1-d_res_2	$1-o_pog_2	±19	$1-d_gran_2	±5

AT2503B (γ) s/n: $bd_num2$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	The range of relative response variation limits error Rwi, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$2-d0		$2-av-bg	—	—	—	—
0.80 µSv/h	$2-n0	$2-d0		$2-av0	$2-o0	±20	$2-q0	±20
4.0 mSv/h	$2-n1	$2-d1		$2-av1	$2-o1	±19	$2-q1	±15
80.0 mSv/h	$2-n2	$2-d2		$2-av2	$2-o2	±19	$2-q2	±15
400 mSv/h	$2-n3	$2-d3		$2-av3	$2-o3	±19	$2-q3	±15
800 mSv/h	$2-n4	$2-d4		$2-av4	$2-o4	±19	$2-q4	±15

AT2503B (γ) s/n: $bd_num2$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$2-dn_1 / $2-dd_1	$2-d_dev_1	$2-dose_1	$2-d_res_1	$2-o_pog_1	±20	$2-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$2-dn_2 / $2-dd_2	$2-d_dev_2	$2-dose_2	$2-d_res_2	$2-o_pog_2	±19	$2-d_gran_2	±5

AT2503B (γ) s/n: $bd_num3$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$3-d0		$3-av-bg	—	—	—	—
0.80 µSv/h	$3-n0	$3-d0		$3-av0	$3-o0	±20	$3-q0	±20
4.0 mSv/h	$3-n1	$3-d1		$3-av1	$3-o1	±19	$3-q1	±15
80.0 mSv/h	$3-n2	$3-d2		$3-av2	$3-o2	±19	$3-q2	±15
400 mSv/h	$3-n3	$3-d3		$3-av3	$3-o3	±19	$3-q3	±15
800 mSv/h	$3-n4	$3-d4		$3-av4	$3-o4	±19	$3-q4	±15

AT2503B (γ) s/n: $bd_num3$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$3-dn_1 / $3-dd_1	$3-d_dev_1	$3-dose_1	$3-d_res_1	$3-o_pog_1	±20	$3-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$3-dn_2 / $3-dd_2	$3-d_dev_2	$3-dose_2	$3-d_res_2	$3-o_pog_2	±19	$3-d_gran_2	±5

AT2503B (γ) s/n: $bd_num4$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$4-d0		$4-av-bg	—	—	—	—
0.80 µSv/h	$4-n0	$4-d0		$4-av0	$4-o0	±20	$4-q0	±20
4.0 mSv/h	$4-n1	$4-d1		$4-av1	$4-o1	±19	$4-q1	±15
80.0 mSv/h	$4-n2	$4-d2		$4-av2	$4-o2	±19	$4-q2	±15
400 mSv/h	$4-n3	$4-d3		$4-av3	$4-o3	±19	$4-q3	±15
800 mSv/h	$4-n4	$4-d4		$4-av4	$4-o4	±19	$4-q4	±15

AT2503B (γ) s/n: $bd_num4$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$4-dn_1 / $4-dd_1	$4-d_dev_1	$4-dose_1	$4-d_res_1	$4-o_pog_1	±20	$4-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$4-dn_2 / $4-dd_2	$4-d_dev_2	$4-dose_2	$4-d_res_2	$4-o_pog_2	±19	$4-d_gran_2	±5

AT2503B (γ) s/n: $bd_num5$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$5-d0		$5-av-bg	—	—	—	—
0.80 µSv/h	$5-n0	$5-d0		$5-av0	$5-o0	±20	$5-q0	±20
4.0 mSv/h	$5-n1	$5-d1		$5-av1	$5-o1	±19	$5-q1	±15
80.0 mSv/h	$5-n2	$5-d2		$5-av2	$5-o2	±19	$5-q2	±15
400 mSv/h	$5-n3	$5-d3		$5-av3	$5-o3	±19	$5-q3	±15
800 mSv/h	$5-n4	$5-d4		$5-av4	$5-o4	±19	$5-q4	±15

AT2503B (γ) s/n: $bd_num5$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$5-dn_1 / $5-dd_1	$5-d_dev_1	$5-dose_1	$5-d_res_1	$5-o_pog_1	±20	$5-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$5-dn_2 / $5-dd_2	$5-d_dev_2	$5-dose_2	$5-d_res_2	$5-o_pog_2	±19	$5-d_gran_2	±5

AT2503B (γ) s/n: $bd_num6$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$6-d0		$6-av-bg	—	—	—	—
0.80 µSv/h	$6-n0	$6-d0		$6-av0	$6-o0	±20	$6-q0	±20
4.0 mSv/h	$6-n1	$6-d1		$6-av1	$6-o1	±19	$6-q1	±15
80.0 mSv/h	$6-n2	$6-d2		$6-av2	$6-o2	±19	$6-q2	±15
400 mSv/h	$6-n3	$6-d3		$6-av3	$6-o3	±19	$6-q3	±15
800 mSv/h	$6-n4	$6-d4		$6-av4	$6-o4	±19	$6-q4	±15

AT2503B (γ) s/n: $bd_num6$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$6-dn_1 / $6-dd_1	$6-d_dev_1	$6-dose_1	$6-d_res_1	$6-o_pog_1	±20	$6-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$6-dn_2 / $6-dd_2	$6-d_dev_2	$6-dose_2	$6-d_res_2	$6-o_pog_2	±19	$6-d_gran_2	±5

AT2503B (γ) s/n: $bd_num7$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$7-d0		$7-av-bg	—	—	—	—
0.80 µSv/h	$7-n0	$7-d0		$7-av0	$7-o0	±20	$7-q0	±20
4.0 mSv/h	$7-n1	$7-d1		$7-av1	$7-o1	±19	$7-q1	±15
80.0 mSv/h	$7-n2	$7-d2		$7-av2	$7-o2	±19	$7-q2	±15
400 mSv/h	$7-n3	$7-d3		$7-av3	$7-o3	±19	$7-q3	±15
800 mSv/h	$7-n4	$7-d4		$7-av4	$7-o4	±19	$7-q4	±15

AT2503B (γ) s/n: $bd_num7$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$7-dn_1 / $7-dd_1	$7-d_dev_1	$7-dose_1	$7-d_res_1	$7-o_pog_1	±20	$7-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$7-dn_2 / $7-dd_2	$7-d_dev_2	$7-dose_2	$7-d_res_2	$7-o_pog_2	±19	$7-d_gran_2	±5

AT2503B (γ) s/n: $bd_num8$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$8-d0		$8-av-bg	—	—	—	—
0.80 µSv/h	$8-n0	$8-d0		$8-av0	$8-o0	±20	$8-q0	±20
4.0 mSv/h	$8-n1	$8-d1		$8-av1	$8-o1	±19	$8-q1	±15
80.0 mSv/h	$8-n2	$8-d2		$8-av2	$8-o2	±19	$8-q2	±15
400 mSv/h	$8-n3	$8-d3		$8-av3	$8-o3	±19	$8-q3	±15
800 mSv/h	$8-n4	$8-d4		$8-av4	$8-o4	±19	$8-q4	±15

AT2503B (γ) s/n: $bd_num8$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$8-dn_1 / $8-dd_1	$8-d_dev_1	$8-dose_1	$8-d_res_1	$8-o_pog_1	±20	$8-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$8-dn_2 / $8-dd_2	$8-d_dev_2	$8-dose_2	$8-d_res_2	$8-o_pog_2	±19	$8-d_gran_2	±5

AT2503B (γ) s/n: $bd_num9$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$9-d0		$9-av-bg	—	—	—	—
0.80 µSv/h	$9-n0	$9-d0		$9-av0	$9-o0	±20	$9-q0	±20
4.0 mSv/h	$9-n1	$9-d1		$9-av1	$9-o1	±19	$9-q1	±15
80.0 mSv/h	$9-n2	$9-d2		$9-av2	$9-o2	±19	$9-q2	±15
400 mSv/h	$9-n3	$9-d3		$9-av3	$9-o3	±19	$9-q3	±15
800 mSv/h	$9-n3	$9-d3		$9-av3	$9-o3	±19	$9-q3	±15

AT2503B (γ) s/n: $bd_num9$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$9-dn_1 / $9-dd_1	$9-d_dev_1	$9-dose_1	$9-d_res_1	$9-o_pog_1	±20	$9-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$9-dn_2 / $9-dd_2	$9-d_dev_2	$9-dose_2	$9-d_res_2	$9-o_pog_2	±19	$9-d_gran_2	±5

AT2503B (γ) s/n: $bd_num10$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$10-d0		$10-av-bg	—	—	—	—
0.80 µSv/h	$10-n0	$10-d0		$10-av0	$10-o0	±20	$10-q0	±20
4.0 mSv/h	$10-n1	$10-d1		$10-av1	$10-o1	±19	$10-q1	±15
80.0 mSv/h	$10-n2	$10-d2		$10-av2	$10-o2	±19	$10-q2	±15
400 mSv/h	$10-n3	$10-d3		$10-av3	$10-o3	±19	$10-q3	±15
800 mSv/h	$10-n4	$10-d4		$10-av4	$10-o4	±19	$10-q4	±15

AT2503B (γ) s/n: $bd_num10$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$10-dn_1 / $10-dd_1	$10-d_dev_1	$10-dose_1	$10-d_res_1	$10-o_pog_1	±20	$10-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$10-dn_2 / $10-dd_2	$10-d_dev_2	$10-dose_2	$10-d_res_2	$10-o_pog_2	±19	$10-d_gran_2	±5

AT2503B (γ) s/n: $bd_num11$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$11-d0		$11-av-bg	—	—	—	—
0.80 µSv/h	$11-n0	$11-d0		$11-av0	$11-o0	±20	$11-q0	±20
4.0 mSv/h	$11-n1	$11-d1		$11-av1	$11-o1	±19	$11-q1	±15
80.0 mSv/h	$11-n2	$11-d2		$11-av2	$11-o2	±19	$11-q2	±15
400 mSv/h	$11-n3	$11-d3		$11-av3	$11-o3	±19	$11-q3	±15
800 mSv/h	$11-n4	$11-d4		$11-av4	$11-o4	±19	$11-q4	±15

AT2503B (γ) s/n: $bd_num11$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$11-dn_1 / $11-dd_1	$11-d_dev_1	$11-dose_1	$11-d_res_1	$11-o_pog_1	±20	$11-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$11-dn_2 / $11-dd_2	$11-d_dev_2	$11-dose_2	$11-d_res_2	$11-o_pog_2	±19	$11-d_gran_2	±5

AT2503B (γ) s/n: $bd_num12$
Dose rate at check point, / Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point, / Ḣ*i(10)	Average value, / Hpi(10)	Relative measure-ment error / θпрi, %	Relative standard deviation, / Si, %	Confidence limit of the intrinsic relative error Δi, %	Limits of intrinsic relative / error, % / not above
Background	—	$12-d0		$12-av-bg	—	—	—	—
0.80 µSv/h	$12-n0	$12-d0		$12-av0	$12-o0	±20	$12-q0	±20
4.0 mSv/h	$12-n1	$12-d1		$12-av1	$12-o1	±19	$12-q1	±15
80.0 mSv/h	$12-n2	$12-d2		$12-av2	$12-o2	±19	$12-q2	±15
400 mSv/h	$12-n3	$12-d3		$12-av3	$12-o3	±19	$12-q3	±15
800 mSv/h	$12-n4	$12-d4		$12-av4	$12-o4	±19	$12-q4	±15

AT2503B (γ) s/n: $bd_num12$
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hp0(10)
4.0 µSv	40 µSv/h / 360s	$12-dn_1 / $12-dd_1	$12-d_dev_1	$12-dose_1	$12-d_res_1	$12-o_pog_1	±20	$12-d_gran_1	±5
4.0 mSv	80 mSv/h / 180s	$12-dn_2 / $12-dd_2	$12-d_dev_2	$12-dose_2	$12-d_res_2	$12-o_pog_2	±19	$12-d_gran_2	±5

Calibrated by:	$calibratedBy
		(signature)
Technical control:	$controlledBy
		(signature)