CALIBRATION CERTIFICATE

Type: Radiation monitor AT1117M	Date of calibration: $date

SN: $serial
Dose rate measurement limits of X-ray and gamma radiation:

PU detector: 1.0 Sv/h – 10.0 mSv/h;

PU detector: 1.0 Sv/h – 100.0 mSv/h;

PU-2 detector: 1.0 mSv/h – 100.0 mSv/h;

PU-2 detector: 1.0 mSv/h – 10.0 mSv/h;

PU-4 detector: 0.3 Sv/h – 0.1 Sv/h;

BDKG-01: 0.1 Sv/h – 10.0 Sv/h;

BDKG-03: 30.0 nSv/h – 0.3 mSv/h;

BDKG-04: 50.0 nSv/h – 10.0 Sv/h;

BDKG-05: 30.0 nSv/h – 0.3 mSv/h;

BDKG-11: 10 nSv/h – 0.1 mSv/h;

BDKG-17: 1.0 mSv/h – 100 Sv/h;

BDKG-24: 20 nSv/h – 1.0 Sv/h;

BDKG-32: 20.0 nSv/h – 0.5 Sv/h;

BDKR-01: 0.05 mSv/h – 100 mSv/h;

BDPS-02:  0.10 Sv/h – 30 mSv/h

Flux density measurement limits of α particles:

BDPA-01:  0.1 – 105 min-1 cm-2;

BDPA-02:  0.05 – 5104 min-1cm -2;

BDPA-03:  0.05 – 2104 min-1cm -2;

BDPS-02:  2.4 – 106 min-1cm-2;

Flux density measurement limits of β particles:

BDPB-01:  1 –5105 min-1cm-2;

BDPB-02:  0.5 – 1.5105 min-1cm-2;

BDPB-03:  0.5 – 0.5105 min-1cm-2;

BDPS-03:  6.0 – 106 min-1cm-2;

Neutron flux density measurement limits:

BDKN-01: 0.1 – 104 s-1cm-2 ;

BDKN-05: 0.1 – 2103 s-1cm-2 ;

Neutron radiation dose rate measurement limits:

BDKN-03: 0.1 mSv/h – 10.0 mSv/h;

X-ray and gamma radiation air kerma rate measurement limits:

BDKG-30: 0.02 Gy/h – 1 mGy/h;

03: ± 20 %БД
Measurement error:
PU: ± 20 %;

PU: ± 20 %;

PU-2: ± 20 %;

PU-2: ± 20 %;

PU-4: ± 20 %;

BDKG-01: ± 20 %;

BDKG-03: ± 20 %;

BDKG-04: ± 20 %;

BDKG-05: ± 20 %;

BDKG-11: ± 20%;

BDKG-17: ± 20 %;

BDKG-24: ± 20 %;

BDKG-30: ± 20 %;

BDKG-32: ± 20 %;

BDKN-01: ± 20 %;

BDKN-03: ± 20 %;

BDKN-05: ± 20 %;

BDPA-01: ± 20 %;

BDPA-02: ± 20 %;

BDPA-03: ± 20 %;

BDPB-01: ± 20 %;

BDPB-02: ± 20 %;

BDPB-03: ± 20 %;

BDKR-01: ± 20 %;

BDPS-02:  2,4 – 30 min-1 sm- 2 – ± 30 % /  30 – 106 min-1 sm- 2 – ± 20 % /  6 – 106 min-1 sm- 2 – ± 20 % /  0.10 Sv/h – 30 mSv/h – ± 20 %

Operating conditions:
- Air temperature	+$temp °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	$vlag %
- Gamma radiation background	$bground nSv/h

Calibration means:
- $eng_facility_110;

- $eng_facility_130;

- $eng_facility_140

- $eng_facility_alpha;

- $eng_facility_beta;

Calibration data
PU ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
20 Sv/h	$16-30	$16-40	$16-50	$16-60	$16-70	$16-80	$16-00	$16-10	$16-20	±20
70 Sv/h	$16-31	$16-41	$16-51	$16-61	$16-71	$16-81	$16-01	$16-11	$16-21
0,7 mSv/h	$16-32	$16-42	$16-52	$16-62	$16-72	$16-82	$16-02	$16-12	$16-22
7 mSv/h	$16-33	$16-43	$16-53	$16-63	$16-73	$16-83	$16-03	$16-13	$16-23

PU-2 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
20 Sv/h	$17-30	$17-40	$17-50	$17-60	$17-70	$17-80	$17-00	$17-10	$17-20	±20
70 Sv/h	$17-31	$17-41	$17-51	$17-61	$17-71	$17-81	$17-01	$17-11	$17-21
0,7 mSv/h	$17-32	$17-42	$17-52	$17-62	$17-72	$17-82	$17-02	$17-12	$17-22
7 mSv/h	$17-33	$17-43	$17-53	$17-63	$17-73	$17-83	$17-03	$17-13	$17-23

PU ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
20 Sv/h	$23-30	$23-40	$23-50	$23-60	$23-70	$23-80	$23-00	$23-10	$23-20	±20
70 Sv/h	$23-31	$23-41	$23-51	$23-61	$23-71	$23-81	$23-01	$23-11	$23-21
0,7 mSv/h	$23-32	$23-42	$23-52	$23-62	$23-72	$23-82	$23-02	$23-12	$23-22
7 mSv/h	$23-33	$23-43	$23-53	$23-63	$23-73	$23-83	$23-03	$23-13	$23-23
70 mSv/h	$23-34	$23-44	$23-54	$23-64	$23-74	$23-84	$23-04	$23-14	$23-24

PU-2 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
20 Sv/h	$24-30	$24-40	$24-50	$24-60	$24-70	$24-80	$24-00	$24-10	$24-20	±20
70 Sv/h	$24-31	$24-41	$24-51	$24-61	$24-71	$24-81	$24-01	$24-11	$24-21
0,7 mSv/h	$24-32	$24-42	$24-52	$24-62	$24-72	$24-82	$24-02	$24-12	$24-22
7 mSv/h	$24-33	$24-43	$24-53	$24-63	$24-73	$24-83	$24-03	$24-13	$24-23
70 mSv/h	$24-34	$24-44	$24-54	$24-64	$24-74	$24-84	$24-04	$24-14	$24-24

PU-4 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
70 Sv/h	$18-30	$18-40	$18-50	$18-60	$18-70	$18-80	$18-00	$18-10	$18-20	±20
0,7 mSv/h	$18-31	$18-41	$18-51	$18-61	$18-71	$18-81	$18-01	$18-11	$18-21
7 mSv/h	$18-32	$18-42	$18-52	$18-62	$18-72	$18-82	$18-02	$18-12	$18-22
70 mSv/h	$18-33	$18-43	$18-53	$18-63	$18-73	$18-83	$18-03	$18-13	$18-23

BDKG-01 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,7 Sv/h	$1-30	$1-40	$1-50	$1-60	$1-70	$1-80	$1-00	$1-10	$1-20	±20
7 Sv/h	$1-31	$1-41	$1-51	$1-61	$1-71	$1-81	$1-01	$1-11	$1-21
70 Sv/h	$1-32	$1-42	$1-52	$1-62	$1-72	$1-82	$1-02	$1-12	$1-22
0.7 mSv/h	$1-33	$1-43	$1-53	$1-63	$1-73	$1-83	$1-03	$1-13	$1-23
7,0 mSv/h	$1-34	$1-44	$1-54	$1-64	$1-74	$1-84	$1-04	$1-14	$1-24
70 mSv/h	$1-35	$1-45	$1-55	$1-65	$1-75	$1-85	$1-05	$1-15	$1-25
0,7 Sv/h	$1-36	$1-46	$1-56	$1-66	$1-76	$1-86	$1-06	$1-16	$1-26
7 Sv/h	$1-37	$1-47	$1-57	$1-67	$1-77	$1-87	$1-07	$1-17	$1-27

BDKG-03 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,07 Sv/h	$2-30	$2-40	$2-50	$2-60	$2-70	$2-80	$2-00	$2-10	$2-20	±20
0,7 Sv/h	$2-31	$2-41	$2-51	$2-61	$2-71	$2-81	$2-01	$2-11	$2-21
7 Sv/h	$2-32	$2-42	$2-52	$2-62	$2-72	$2-82	$2-02	$2-12	$2-22
70 Sv/h	$2-33	$2-43	$2-53	$2-63	$2-73	$2-83	$2-03	$2-13	$2-23
240 Sv/h	$2-34	$2-44	$2-54	$2-64	$2-74	$2-84	$2-04	$2-14	$2-24

BDKG-04 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,07 Sv/h	$3-30	$3-40	$3-50	$3-60	$3-70	$3-80	$3-00	$3-10	$3-20	±20
0,7 Sv/h	$3-31	$3-41	$3-51	$3-61	$3-71	$3-81	$3-01	$3-11	$3-21
7 Sv/h	$3-32	$3-42	$3-52	$3-62	$3-72	$3-82	$3-02	$3-12	$3-22
70 Sv/h	$3-33	$3-43	$3-53	$3-63	$3-73	$3-83	$3-03	$3-13	$3-23
0.7 mSv/h	$3-34	$3-44	$3-54	$3-64	$3-74	$3-84	$3-04	$3-14	$3-24
7,0 mSv/h	$3-35	$3-45	$3-55	$3-65	$3-75	$3-85	$3-05	$3-15	$3-25
70 mSv/h	$3-36	$3-46	$3-56	$3-66	$3-76	$3-86	$3-06	$3-16	$3-26
0,7 Sv/h	$3-37	$3-47	$3-57	$3-67	$3-77	$3-87	$3-07	$3-17	$3-27
7 Sv/h	$3-38	$3-48	$3-58	$3-68	$3-78	$3-88	$3-08	$3-18	$3-28

BDKG-05 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,07 Sv/h	$4-30	$4-40	$4-50	$4-60	$4-70	$4-80	$4-00	$4-10	$4-20	±20
0,7 Sv/h	$4-31	$4-41	$4-51	$4-61	$4-71	$4-81	$4-01	$4-11	$4-21
7 Sv/h	$4-32	$4-42	$4-52	$4-62	$4-72	$4-82	$4-02	$4-12	$4-22
70 Sv/h	$4-33	$4-43	$4-53	$4-63	$4-73	$4-83	$4-03	$4-13	$4-23
240 Sv/h	$4-34	$4-44	$4-54	$4-64	$4-74	$4-84	$4-04	$4-14	$4-24

BDKG-11 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,07 Sv/h	$5-30	$5-40	$5-50	$5-60	$5-70	$5-80	$5-00	$5-10	$5-20	±20
0,7 Sv/h	$5-31	$5-41	$5-51	$5-61	$5-71	$5-81	$5-01	$5-11	$5-21
7 Sv/h	$5-32	$5-42	$5-52	$5-62	$5-72	$5-82	$5-02	$5-12	$5-22
70 Sv/h	$5-33	$5-43	$5-53	$5-63	$5-73	$5-83	$5-03	$5-13	$5-23

BDKG-17 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
7,0 mSv/h	$6-30	$6-40	$6-50	$6-60	$6-70	$6-80	$6-00	$6-10	$6-20	±20
70 mSv/h	$6-31	$6-41	$6-51	$6-61	$6-71	$6-81	$6-01	$6-11	$6-21
0,7 Sv/h	$6-32	$6-42	$6-52	$6-62	$6-72	$6-82	$6-02	$6-12	$6-22
7 Sv/h	$6-33	$6-43	$6-53	$6-63	$6-73	$6-83	$6-03	$6-13	$6-23
40 Sv/h	$6-34	$6-44	$6-54	$6-64	$6-74	$6-84	$6-04	$6-14	$6-24

BDKG-24 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,07 Sv/h	$7-30	$7-40	$7-50	$7-60	$7-70	$7-80	$7-00	$7-10	$7-20	±20
0,7 Sv/h	$7-31	$7-41	$7-51	$7-61	$7-71	$7-81	$7-01	$7-11	$7-21
7 Sv/h	$7-32	$7-42	$7-52	$7-62	$7-72	$7-82	$7-02	$7-12	$7-22
70 Sv/h	$7-33	$7-43	$7-53	$7-63	$7-73	$7-83	$7-03	$7-13	$7-23
0.7 mSv/h	$7-34	$7-44	$7-54	$7-64	$7-74	$7-84	$7-04	$7-14	$7-24
7,0 mSv/h	$7-35	$7-45	$7-55	$7-65	$7-75	$7-85	$7-05	$7-15	$7-25
70 mSv/h	$7-36	$7-46	$7-56	$7-66	$7-76	$7-86	$7-06	$7-16	$7-26
0,7 Sv/h	$7-37	$7-47	$7-57	$7-67	$7-77	$7-87	$7-07	$7-17	$7-27

BDKG-30 ()
Kerma rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Kerma rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,03 Gy/h	$8-30	$8-40	$8-50	$8-60	$8-70	$8-80	$8-00	$8-10	$8-20	±20
0,7 Gy/h	$8-31	$8-41	$8-51	$8-61	$8-71	$8-81	$8-01	$8-11	$8-21
7 Gy/h	$8-32	$8-42	$8-52	$8-62	$8-72	$8-82	$8-02	$8-12	$8-22
70 Gy/h	$8-33	$8-43	$8-53	$8-63	$8-73	$8-83	$8-03	$8-13	$8-23
0.7 mGy/h	$8-34	$8-44	$8-54	$8-64	$8-74	$8-84	$8-04	$8-14	$8-24
7,0 mGy/h	$8-35	$8-45	$8-55	$8-65	$8-75	$8-85	$8-05	$8-15	$8-25
70 mGy/h	$8-36	$8-46	$8-56	$8-66	$8-76	$8-86	$8-06	$8-16	$8-26
0,7 Gy/h	$8-37	$8-47	$8-57	$8-67	$8-77	$8-87	$8-07	$8-17	$8-27

BDKG-32 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,07 Sv/h	$9-30	$9-40	$9-50	$9-60	$9-70	$9-80	$9-00	$9-10	$9-20	±20
0,7 Sv/h	$9-31	$9-41	$9-51	$9-61	$9-71	$9-81	$9-01	$9-11	$9-21
7 Sv/h	$9-32	$9-42	$9-52	$9-62	$9-72	$9-82	$9-02	$9-12	$9-22
70 Sv/h	$9-33	$9-43	$9-53	$9-63	$9-73	$9-83	$9-03	$9-13	$9-23
0.7 mSv/h	$9-34	$9-44	$9-54	$9-64	$9-74	$9-84	$9-04	$9-14	$9-24
7,0 mSv/h	$9-35	$9-45	$9-55	$9-65	$9-75	$9-85	$9-05	$9-15	$9-25
70 mSv/h	$9-36	$9-46	$9-56	$9-66	$9-76	$9-86	$9-06	$9-16	$9-26
0,35 Sv/h	$9-37	$9-47	$9-57	$9-67	$9-77	$9-87	$9-07	$9-17	$9-27
0,5 Sv/h	$9-38	$9-48	$9-58	$9-68	$9-78	$9-88	$9-08	$9-18	$9-28

BDKN-01 (n)
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=6 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,965 / 0i=986 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$10-61 / $10-71 / $10-81	$10-62 / $10-72 / $10-82	$10-63 / $10-73 / $10-83	$10-64 / $10-74 / $10-84
Average readings, i, s-1cm-2	$10-00	$10-01	$10-02	$10-03	$10-04
Result of flux density measurement at control point / npi, s-1cm-2		$10-n21	$10-n22	$10-n23	$10-n24
Intrinsic relative error of measurement, θпрi, %		$10-n31	$10-n32	$10-n33	$10-n34
Confidence limit of the intrinsic relative error Δi, %		$10-n41	$10-n42	$10-n43	$10-n44
Limite of intrinsic relative error , % not above		±20

BDKN-03 (n)
Dose rate at check point Ḣoi(10)	Bacground / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,922 / H*oi(10)=9 / 1-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,922 / H*oi(10)=80 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,980 / H*oi(10)=330 / 2·102-103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,995 / H*oi(10)=1460 / 2·103-104
Readings at check point / H*i(10) , ±δст, %		$11-61 / $11-71 / $11-81	$11-62 / $11-72 / $11-82	$11-63 / $11-73 / $11-83	$11-64 / $11-74 / $11-84
Average readings, H*i(10)	$11-00	$11-01	$11-02	$11-03	$11-04
Dose rate measurement result at control point / H*прi(10)		$11-n21	$11-n22	$11-n23	$11-n24
Intrinsic relative error of measurement, θпрi, %		$11-n31	$11-n32	$11-n33	$11-n34
Confidence limit of the intrinsic relative error Δi, %		$11-n41	$11-n42	$11-n43	$11-n44
Limite of intrinsic relative error , % not above		±20

BDKN-05 (n)
Измеренное значение скорости счета фона, с-1					Условие соответствия по ТУ nф,, имп./с, не более
Результаты измерения			Среднее значение nф
$27-10			$27-20			0,23
$27-11
$27-12
№ источника	Расстояние до источника r0, см	Плотность потока в / i-ой контроль-ной точке, φoi, с-1·см-2	Значение коэффици-ента b(r0)	Значение коэффици-ента К	Показания прибора в контроль-ной точке, N, с-1	Среднее значение показаний N, с-1	Чувстви-тельность БД, S, имп.·см2/ нейтр.	Условие соответ-ствия по ТУ S,, имп·см2, не менее
10	100	2,5 – 500	0,866	0,785	$27-50	$27-60	$27-70	8,0
					$27-51
					$27-52

BDPA-01 (α)
Flux density in test point 0i, / min-1·cm-2	Radiation source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$12-90	1179	$s_bg	$12-60	$12-70	$12-80	$12-00	$12-10	$12-20	±20
$12-91	1180	---	$12-61	$12-71	$12-81	$12-01	$12-11	$12-21
$12-92	1181	---	$12-62	$12-72	$12-82	$12-02	$12-12	$12-22
$12-93	1182	---	$12-63	$12-73	$12-83	$12-03	$12-13	$12-23

BDPA-02 (α)
Flux density in test point 0i, / min-1·cm-2	Radiation source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$13-90	0651	$s_bg	$13-60	$13-70	$13-80	$13-00	$13-10	$13-20	±20
$13-91	0652	---	$13-61	$13-71	$13-81	$13-01	$13-11	$13-21
$13-92	0653	---	$13-62	$13-72	$13-82	$13-02	$13-12	$13-22
$13-93	0655	---	$13-63	$13-73	$13-83	$13-03	$13-13	$13-23

BDPA-03 (α)
Flux density in test point 0i, / min-1·cm-2	Radiation source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$25-90	0651	$s_bg	$25-60	$25-70	$25-80	$25-00	$25-10	$25-20	±20
$25-91	0652	---	$25-61	$25-71	$25-81	$25-01	$25-11	$25-21
$25-92	0653	---	$25-62	$25-72	$25-82	$25-02	$25-12	$25-22
$25-93	0655	---	$25-63	$25-73	$25-83	$25-03	$25-13	$25-23

BDPB-01 (β)
Flux density in test point 0i, / min-1·cm-2	Radiation source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$14-90	$14-30	$b_bg	$14-60	$14-70	$14-80	$14-00	$14-10	$14-20	±20
$14-91	$14-31	---	$14-61	$14-71	$14-81	$14-01	$14-11	$14-21
$14-92	$14-32	---	$14-62	$14-72	$14-82	$14-02	$14-12	$14-22
$14-93	$14-33	---	$14-63	$14-73	$14-83	$14-03	$14-13	$14-23
$14-94	$14-34	---	$14-64	$14-74	$14-84	$14-04	$14-14	$14-24

BDPB-02 (β)
Flux density in test point 0i, / min-1·cm-2	Radiation source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$15-90	$15-30	$b2_bg	$15-60	$15-70	$15-80	$15-00	$15-10	$15-20	±20
$15-91	$15-31	---	$15-61	$15-71	$15-81	$15-01	$15-11	$15-21
$15-92	$15-32	---	$15-62	$15-72	$15-82	$15-02	$15-12	$15-22
$15-93	$15-33	---	$15-63	$15-73	$15-83	$15-03	$15-13	$15-23
$15-94	$15-34	---	$15-64	$15-74	$15-84	$15-04	$15-14	$15-24

BDPB-03 (β)
Flux density in test point 0i, / min-1·cm-2	Radiation source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$26-90	0663	$b2_bg	$26-60	$26-70	$26-80	$26-00	$26-10	$26-20	±20
$26-91	0664	---	$26-61	$26-71	$26-81	$26-01	$26-11	$26-21
$26-92	0665	---	$26-62	$26-72	$26-82	$26-02	$26-12	$26-22
$26-93	0666	---	$26-63	$26-73	$26-83	$26-03	$26-13	$26-23
$26-94	0667	---	$26-64	$26-74	$26-84	$26-04	$26-14	$26-24

BDKR-01 ()
Dose rate at check point Ḣoi(0.07)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(0.07)			Average value, Hi(0.07)
				H1	H2	H3
0,07 Sv/h	$19-30	$19-40	$19-50	$19-60	$19-70	$19-80	$19-00	$19-10	$19-20	±20
0,7 Sv/h	$19-31	$19-41	$19-51	$19-61	$19-71	$19-81	$19-01	$19-11	$19-21
7 Sv/h	$19-32	$19-42	$19-52	$19-62	$19-72	$19-82	$19-02	$19-12	$19-22
70 Sv/h	$19-33	$19-43	$19-53	$19-63	$19-73	$19-83	$19-03	$19-13	$19-23

BDPS-02 (α)
Flux density in test point 0i, / min-1·cm-2	Radiation source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$20-90	1179	---	$20-60	$20-70	$20-80	$20-00	$20-10	$20-20	±20
$20-91	1180	---	$20-61	$20-71	$20-81	$20-01	$20-11	$20-21
$20-92	1181	---	$20-62	$20-72	$20-82	$20-02	$20-12	$20-22
$20-93	1182	---	$20-63	$20-73	$20-83	$20-03	$20-13	$20-23
$20-94	1183	---	$20-64	$20-74	$20-84	$20-04	$20-14	$20-24

BDPS-02 (β)
Flux density in test point 0i, / min-1·cm-2	Radiation source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$21-90	$21-30	---	$21-60	$21-70	$21-80	$21-00	$21-10	$21-20	±20
$21-91	$21-31	---	$21-61	$21-71	$21-81	$21-01	$21-11	$21-21
$21-92	$21-32	---	$21-62	$21-72	$21-82	$21-02	$21-12	$21-22
$21-93	$21-33	---	$21-63	$21-73	$21-83	$21-03	$21-13	$21-23
$21-94	$21-34	---	$21-64	$21-74	$21-84	$21-04	$21-14	$21-24

BDPS-02 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, / Hi(10)
				H1	H2	H3
0,7 Sv/h	$22-30	$22-40	$22-50	$22-60	$22-70	$22-80	$22-00	$22-10	$22-20	±20
7 Sv/h	$22-31	$22-41	$22-51	$22-61	$22-71	$22-81	$22-01	$22-11	$22-21
70 Sv/h	$22-32	$22-42	$22-52	$22-62	$22-72	$22-82	$22-02	$22-12	$22-22
0.7 mSv/h	$22-33	$22-43	$22-53	$22-63	$22-73	$22-83	$22-03	$22-13	$22-23
7,0 mSv/h	$22-34	$22-44	$22-54	$22-64	$22-74	$22-84	$22-04	$22-14	$22-24
20 mSv/h	$22-35	$22-45	$22-55	$22-65	$22-75	$22-85	$22-05	$22-15	$22-25

Calibrated by:	$calibratedBy
		(signature)
Technical control:	$controlledBy
		(signature)