CALIBRATION CERTIFICATE
AT2327 Alarm Dosimeter	Date of calibration: $date

SN: $serial

Type: BDKN-05 detection unit

Sensitivity to Pu-α-Be neutron source radiation:	≥8 impulses·cm2/neutron

Measurement error:	±20%

Operating conditions:
- Air temperature	+$temp °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	$vlag %
- Gamma radiation background	$bground nSv/h

Calibration means:
$eng_facility_140;·

Calibration data
a)	Intrinsic background level of spectrometer with neutron radiation detection unit

BDKN-05 (n) s/n: $bd_num1$
Intrinsic background level, s-1		Maximum background level, s-1
Readings at check point	Average value nф
$1-10	$1-20	0,23
$1-11
$1-12

b) Sensitivity to neutron radiation / BDKN-05 (n) s/n: $bd_num1$
Flux / density at check point φ0i , s-1·cm-2	Radiation source number	Distance to source, / R, сm	Value of coefficient К b(r0)	Value of coefficient К	Measurement at check point, N, с-1	Average value, / N, с-1	Sensitivity / in the direct beam Pu-Be neutron source, impulses·cm2/neutron	Minimal detectable sensitivity, impulses·cm2/neutron
Background	–	–	–	–	$1-30	$1-40	–	–
					$1-31
					$1-32
2,5 – 500	10	100	0,866	0,785	$1-50	$1-60	$1-70	8,0
					$1-51
					$1-52

Calibrated by:	$calibratedBy
		(signature)
Technical control:	$controlledBy
		(signature)