CALIBRATION CERTIFICATE
Type: Radiation monitor $type	Date of calibration: 29.11.2021.

S/N: 1234

Measurement limits of ambient x-ray gamma radiation:

АТ1125А: γ 30 nSv/h – 100 mSv/h;

03: ± 20 %БД
Measurement error:
АТ1125А: ± 15 %;

Operating conditions:
- Air temperature	+22.0 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	74.0 %
- Gamma radiation background	95.0 nSv/h

Calibration means:
- standard dosimetry facility AT-110, N 013, Cs-137, error 5%, the Certificate of Compliance N210/1426-2018 on 22.10.2018 issued by FGUP «D.I.Mendeleyev VNIIM», St. Petersburg, Russia);

- standard dosimetry facility AT-130, N 015, Cs-137-137, error 4 %, the Certificate of Compliance N210/ 1427-2018 on 23.10.2018 issued by FGUP «D.I.Mendeleyev VNIIM», St. Petersburg, Russia);

Calibration data / Radiation monitor 1125A s/n 1234
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,24 µSv/h	263	106.1	95.0	0.24	0.25	0.23	0.24	0.00	5.50	±15
0,7 µSv/h	0НА	230.1	95.0	0.71	0.68	0.71	0.70	0.00	5.50
7 µSv/h	0НА	73.9	—	6.95	7.13	6.92	7.00	0.00	5.50
70 µSv/h	9ХК	159.8	—	69.6	69.3	71.1	70.0	0.00	5.50
240 µSv/h	9ХК	86.8	—	246	243	231	240	0.00	5.50
0,7 mSv/h	9ХК	51.3	—	0.71	0.72	0.67	0.70	0.00	5.50
7 mSv/h	043	345.5	—	7.02	6.82	7.15	7.00	0.00	4.40
70 mSv/h	043	111.0	—	68.1	69.0	72.9	70.0	0.00	4.40

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)

Type: Radiation monitor $type	Date of calibration: 29.11.2021.

S/N: 1234

Measurement limits of ambient x-ray gamma radiation:

АТ1125А: γ 30 nSv/h – 100 mSv/h;

03: ± 20 %БД
Measurement error:
АТ1125А: ± 15 %;

Operating conditions:
- Air temperature	+22.0 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	74.0 %
- Gamma radiation background	95.0 nSv/h

Calibration means:
- standard dosimetry facility AT-110, N 013, Cs-137, error 5%, the Certificate of Compliance N210/1426-2018 on 22.10.2018 issued by FGUP «D.I.Mendeleyev VNIIM», St. Petersburg, Russia);

- standard dosimetry facility AT-130, N 015, Cs-137-137, error 4 %, the Certificate of Compliance N210/ 1427-2018 on 23.10.2018 issued by FGUP «D.I.Mendeleyev VNIIM», St. Petersburg, Russia);

Calibration data / Radiation monitor 1125A s/n 1234
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,24 µSv/h	263	106.1	95.0	0.24	0.25	0.23	0.24	0.00	5.50	±15
0,7 µSv/h	0НА	230.1	95.0	0.71	0.68	0.71	0.70	0.00	5.50
7 µSv/h	0НА	73.9	—	6.95	7.13	6.92	7.00	0.00	5.50
70 µSv/h	9ХК	159.8	—	69.6	69.3	71.1	70.0	0.00	5.50
240 µSv/h	9ХК	86.8	—	246	243	231	240	0.00	5.50
0,7 mSv/h	9ХК	51.3	—	0.71	0.72	0.67	0.70	0.00	5.50
7 mSv/h	043	345.5	—	7.02	6.82	7.15	7.00	0.00	4.40
70 mSv/h	043	111.0	—	68.1	69.0	72.9	70.0	0.00	4.40

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)