CALIBRATION CERTIFICATE

Type: Personal dosimeter AT2503	Date of calibration: $date

S/n: $allSerial

Measurement range for:
- individual dose equivalent:	1 µSv – 10 Sv
- individual dose equivalent rate:	0.1 µSv/h – 0.5 Sv/h
Measurement error:
- dose response variation limits	±(15+UrelD)%
- dose rate response variation limits	±(15+UrelD)%
where UrelD is the relative expanded uncertainty of the conventional value of a dose as a percentage
- dose coefficient of variation (H<11H0)	±(16-H/H0)%
- dose coefficient of variation (H226511H0)	±5%
- dose rate coefficient of variation (Ḣ<10 µSv/h)	±20%
- dose rate coefficient of variation (10µSv/h⩽Ḣ<60µSv/h)	±(21-Ḣ/10µSv/h)%
- dose rate coefficient of variation (Ḣ⩾60µSv/h)	±15%
where H0 is the lower limit of dose measurement

Operating conditions:
- air temperature	+$temp °C
- relative humidity	$vlag %
- gamma radiation background	$bground nSv/h

Calibration means:
- $eng_facility_110;
- $eng_facility_130;
Calibration data:
AT2503 (γ) s/n: $bd_num1$
Dose rate at check point, / Ḣpi(10)	Radia-tion source number	Distance to source, / R, сm	Dose rate measurement / at check point, / Ḣpi(10)	Average value, / Ḣpi(10)	Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, / %
Background	—	—		$1-av-bg	—	—	—	—
0.80 µSv/h	$1-n0	$1-d0		$1-av0	$1-o0	±20	$1-q0	±33.8
4.0 mSv/h	$1-n1	$1-d1		$1-av1	$1-o1	±19	$1-q1	±15.0
80.0 mSv/h	$1-n2	$1-d2		$1-av2	$1-o2	±19	$1-q2	±15.4
400 mSv/h	$1-n3	$1-d3		$1-av3	$1-o3	±19	$1-q3	±17.0

AT2503 (γ) s/n: ##bd_num1##
Actual dose value, / Hp0(10)	Dose rate, / Ḣoi(10) / / t, s	Radiation source / / d, cm	Measured dose value			Response variation / Rwi, %	Response variation limits, Rwi, %	Coefficient of variation vi, %	Coefficient / of variation limits, %
			Hpi(10)	Hpi(10)meas.	Hpi(10)meas.-Hpi(10)
4.0 µSv	40 µSv/h / 360s	##1-dn_1 / ##1-dd_1	##1-d_dev_1	##1-dose_1	##1-d_res_1	##1-o_pog_1	±20	##1-d_gran_1	±15.0
4.0 mSv	80 mSv/h / 180s	##1-dn_2 / ##1-dd_2	##1-d_dev_2	##1-dose_2	##1-d_res_2	##1-o_pog_2	±19	##1-d_gran_2	±15.4

Calibrated by:	$calibratedBy
		(signature)
Technical control:	$controlledBy
		(signature)