CALIBRATION CERTIFICATE
Date of calibration:	$date
Type:	Radiation Monitor AT1117M
S/N:	$serial

Dose rate measurement limits of X-ray and gamma radiation, Ḣ*(10):

PU detector: 1.0 µSv/h – 10.0 mSv/h, measurement error: ± 20 %

PU detector: 1.0 µSv/h – 100.0 mSv/h, measurement error: ± 20 %

PU-2 detector: 1.0 µSv/h – 100.0 mSv/h, measurement error: ± 20 %

PU-2 detector: 1.0 µSv/h – 10.0 mSv/h, measurement error: ± 20 %

PU-4 detector: 1.0 µSv/h – 100.0 mSv/h, measurement error: ± 20 %

PU-5 detector: 1.0 µSv/h – 100.0 mSv/h, measurement error: ± 20 %

BDKG-01: 0.1 µSv/h – 10.0 Sv/h, measurement error: ± 20 %

BDKG-03: 30.0 nSv/h – 300 µSv/h, measurement error: ± 20 %

BDKG-04: 50.0 nSv/h – 10.0 Sv/h, measurement error: ± 20 %

BDKG-05: 30.0 nSv/h – 300 µSv/h, measurement error: ± 20 %

BDKG-11: 30.0 nSv/h – 100 µSv/h, measurement error: ± 20 %

BDKG-17: 1.0 mSv/h – 100 Sv/h, measurement error: ± 20 %

BDKG-24: 30 nSv/h – 1.0 Sv/h, measurement error: ± 20 %

BDKG-32: 30.0 nSv/h – 0.5 Sv/h, measurement error: ± 20 %

BDPS-02: γ 0.10 µSv/h – 30 mSv/h, measurement error: ± 20 %

Dose rate measurement limits of X-ray and gamma radiation, Ḣ*(0.07):

BDKR-01: 0.05 µSv/h – 100 µSv/h

Flux density measurement limits of α particles, φ:

BDPA-01: 0.1 – 105 min-1⋅ cm-2

BDPA-02: 0.05 – 5⋅104 min-1⋅cm -2

BDPA-03: 0.05 – 2⋅104 min-1⋅cm -2

BDPS-02: 2.4 – 106 min-1⋅cm-2

Flux density measurement limits of β particles, φ:

BDPB-01: 1 – 5⋅105 min-1⋅cm-2;

BDPB-02: 0.5 – 1.5⋅105 min-1⋅cm-2

BDPB-03: 0.5 – 0.5⋅105 min-1⋅cm-2

BDPS-02: 6.0 – 106 min-1⋅cm-2

Neutron flux density measurement limits, φ:

BDKN-01: 0.1 – 104 s-1⋅cm-2

BDKN-05: 0.1 – 2⋅103 s-1⋅cm-2

Neutron radiation dose rate measurement limits, Ḣ*(10):

BDKN-03: 0.1 µSv/h – 10.0 mSv/h

BDKN-06: 0.1 µSv/h – 30.0 mSv/h

X-ray and gamma radiation air kerma rate measurement limits, {dot { Ka }} :

BDKG-30: 30 nGy/h – 1 Gy/h;

03: ± 20 %БД
Measurement error:
PU: ± 20 %

PU: ± 20 %

PU-2: ± 20 %

PU-2: ± 20 %

PU-4: ± 20 %

PU-5: ± 20 %

BDKG-01: ± 20 %

BDKG-03: ± 20 %

BDKG-04: ± 20 %

BDKG-05: ± 20 %

BDKG-11: ± 20 %

BDKG-17: ± 20 %

BDKG-24: ± 20 %

BDKG-30: ± 20 %

BDKG-32: ± 20 %

BDKN-01: ± 20 %

BDKN-03: ± 20 %

BDKN-05: ± 20 %

BDKN-06: ± 20 %

BDPA-01: ± 20 %

BDPA-02: ± 20 %

BDPA-03: ± 20 %

BDPB-01: ± 20 %

BDPB-02: ± 20 %

BDPB-03: ± 20 %

BDKR-01: ± 20 %

BDPS-02:	α: 2,4 – 30 min-1⋅ sm- 2 – ± 30 % / α: 30 – 106 min-1⋅ sm- 2 – ± 20 % / β: 6 – 106 min-1⋅ sm- 2 – ± 20 % / γ: 0.10 µSv/h – 30 mSv/h – ± 20 %

Operating conditions:
- Air temperature	+$temp °C
- Atmospheric pressure	$preasure kPa
- Relative humidity	$vlag %
- Gamma radiation background	$bground nSv/h

Calibration means:
- $eng_facility_110$;

- $eng_facility_130$;

- $eng_facility_140$;

- $eng_facility_alpha$;

- $eng_facility_beta$;

Calibration data

PU (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
20 µSv/h	$16-00 µSv/h
70 µSv/h	$16-01 µSv/h
0,7 mSv/h	$16-02 mSv/h
7 mSv/h	$16-03 mSv/h

PU-2 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
20 µSv/h	$17-00 µSv/h
70 µSv/h	$17-01 µSv/h
0,7 mSv/h	$17-02 mSv/h
7 mSv/h	$17-03 mSv/h

PU (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
7 µSv/h	$23-00 µSv/h
70 µSv/h	$23-01 µSv/h
0,7 mSv/h	$23-02 mSv/h
7 mSv/h	$23-03 mSv/h
70 mSv/h	$23-04 mSv/h

PU-2 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
7 µSv/h	$24-00 µSv/h
70 µSv/h	$24-01 µSv/h
0,7 mSv/h	$24-02 mSv/h
7 mSv/h	$24-03 mSv/h
70 mSv/h	$24-04 mSv/h

PU-4 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
7 µSv/h	$18-00 µSv/h
70 µSv/h	$18-01 µSv/h
0,7 mSv/h	$18-02 mSv/h
7 mSv/h	$18-03 mSv/h
70 mSv/h	$18-04 mSv/h

PU-5 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
20 µSv/h	$29-00 µSv/h
70 µSv/h	$29-01 µSv/h
0,7 mSv/h	$29-02 mSv/h
7 mSv/h	$29-03 mSv/h
70 mSv/h	$29-04 mSv/h

BDKG-01 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
0,7 µSv/h	$1-00 µSv/h
7 µSv/h	$1-01 µSv/h
70 µSv/h	$1-02 µSv/h
0,7 mSv/h	$1-03 mSv/h
7,0 mSv/h	$1-04 mSv/h
70 mSv/h	$1-05 mSv/h
0,7 Sv/h	$1-06 Sv/h
7 Sv/h	$1-07 Sv/h

BDKG-03 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
0,07 µSv/h	$2-00 µSv/h
0,7 µSv/h	$2-01 µSv/h
7 µSv/h	$2-02 µSv/h
70 µSv/h	$2-03 µSv/h
240 µSv/h	$2-04 µSv/h

BDKG-04 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
0,07 µSv/h	$3-00 µSv/h
0,7 µSv/h	$3-01 µSv/h
7 µSv/h	$3-02 µSv/h
70 µSv/h	$3-03 µSv/h
0,7 mSv/h	$3-04 mSv/h
7,0 mSv/h	$3-05 mSv/h
70 mSv/h	$3-06 mSv/h
0,7 Sv/h	$3-07 Sv/h
7 Sv/h	$3-08 Sv/h

BDKG-05 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
0,07 µSv/h	$4-00 µSv/h
0,7 µSv/h	$4-01 µSv/h
7 µSv/h	$4-02 µSv/h
70 µSv/h	$4-03 µSv/h
240 µSv/h	$4-04 µSv/h

BDKG-11 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
0,07 µSv/h	$5-00 µSv/h
0,7 µSv/h	$5-01 µSv/h
7 µSv/h	$5-02 µSv/h
70 µSv/h	$5-03 µSv/h

BDKG-17 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
7,0 mSv/h	$6-00 mSv/h
70 mSv/h	$6-01 mSv/h
0,7 Sv/h	$6-02 Sv/h
7 Sv/h	$6-03 Sv/h
40 Sv/h	$6-04 Sv/h

BDKG-24 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
0,07 µSv/h	$7-00 µSv/h
0,7 µSv/h	$7-01 µSv/h
7 µSv/h	$7-02 µSv/h
70 µSv/h	$7-03 µSv/h
0,7 mSv/h	$7-04 mSv/h
7,0 mSv/h	$7-05 mSv/h
70 mSv/h	$7-06 mSv/h
0,7 Sv/h	$7-07 Sv/h

BDKG-30 (γ)
Air kerma rate at check point, / {dot {Ka}} rsub {0i} rsup {"*"}	Calibration data, / {bar {dot {Ka}}} rsub {i} rsup {"*"}
0,07 µGy/h	$8-00 µGy/h
0,7 µGy/h	$8-01 µGy/h
7 µGy/h	$8-02 µGy/h
70 µGy/h	$8-03 µGy/h
0,7 mGy/h	$8-04 mGy/h
7,0 mGy/h	$8-05 mGy/h
70 mGy/h	$8-06 mGy/h
0,7 Gy/h	$8-07 Gy/h

BDKG-32 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
0,07 µSv/h	$9-00 µSv/h
0,7 µSv/h	$9-01 µSv/h
7 µSv/h	$9-02 µSv/h
70 µSv/h	$9-03 µSv/h
0,7 mSv/h	$9-04 mSv/h
7,0 mSv/h	$9-05 mSv/h
70 mSv/h	$9-06 mSv/h
0,4 Sv/h	$9-07 Sv/h

BDKN-01 (n)
Flux density at check point φ0i , s-1⋅cm-2	Calibration data, / s-1⋅cm-2
B(R)=$10-br1 / φ0i=$10-fi1	$10-01
B(R)=$10-br2 / φ0i=$10-fi2	$10-02
B(R)=$10-br3 / φ0i=$10-fi3	$10-03
B(R)=$10-br4 / φ0i=$10-fi4	$10-04

BDKN-03 (n)
Dose rate at check point, / Ḣoi(10)	Calibration data, / {bar {dot {H}}} rsub {i} (10)
B(R)=$11-br1 / H*oi(10)=$11-fi1	$11-01
B(R)=$11-br2 / H*oi(10)=$11-fi2	$11-02
B(R)=$11-br3 / H*oi(10)=$11-fi3	$11-03
B(R)=$11-br4 / H*oi(10)=$11-fi4	$11-04

BDKN-06 (n)
Dose rate at check point, / Ḣoi(10)	Calibration data, / {bar {dot {H}}} rsub {i} (10)
B(R)=$28-br1 / H*oi(10)=$28-fi1	$28-01
B(R)=$28-br2 / H*oi(10)=$28-fi2	$28-02
B(R)=$28-br3 / H*oi(10)=$28-fi3	$28-03
B(R)=$28-br4 / H*oi(10)=$28-fi4	$28-04

BDKN-05 (n)
Flux density at check point φ0i , s-1⋅cm-2	Calibration data, / s-1⋅cm-2
B(R)=$27-br1 / φ0i=$27-fi1	$27-01
B(R)=$27-br2 / φ0i=$27-fi2	$27-02
B(R)=$27-br3 / φ0i=$27-fi3	$27-03
B(R)=$27-br4 / φ0i=$27-fi4	$27-04

BDPA-01 (α)
Flux density in test point φ0i, / min-1·cm-2	Calibration data, / min-1·cm-2
$12-90	$12-00
$12-91	$12-01
$12-92	$12-02
$12-93	$12-03

BDPA-02 (α)
Flux density in test point φ0i, / min-1·cm-2	Calibration data, / min-1·cm-2
$13-90	$13-00
$13-91	$13-01
$13-92	$13-02
$13-93	$13-03

BDPA-03 (α)
Flux density in test point φ0i, / min-1·cm-2	Calibration data, / min-1·cm-2
$25-90	$25-00
$25-91	$25-01
$25-92	$25-02
$25-93	$25-03

BDPB-01 (β)
Flux density in test point φ0i, / min-1·cm-2	Calibration data, / min-1·cm-2
$14-90	$14-00
$14-91	$14-01
$14-92	$14-02
$14-93	$14-03
$14-94	$14-04

BDPB-02 (β)
Flux density in test point φ0i, / min-1·cm-2	Calibration data, / min-1·cm-2
$15-90	$15-00
$15-91	$15-01
$15-92	$15-02
$15-93	$15-03
$15-94	$15-04

BDPB-03 (β)
Flux density in test point φ0i, / min-1·cm-2	Calibration data, / min-1·cm-2
$26-90	$26-00
$26-91	$26-01
$26-92	$26-02
$26-93	$26-03
$26-94	$26-04

BDKR-01 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (0.07)	Calibration data, / {bar {dot {H}}} rsub {i} (0.07)
0,07 µSv/h	$19-00 µSv/h
0,7 µSv/h	$19-01 µSv/h
7 µSv/h	$19-02 µSv/h
70 µSv/h	$19-03 µSv/h

BDPS-02 (α)
Flux density in test point φ0i, / min-1·cm-2	Calibration data, / min-1·cm-2
$20-90	$20-00
$20-91	$20-01
$20-92	$20-02
$20-93	$20-03
$20-94	$20-04

BDPS-02 (β)
Flux density in test point φ0i, / min-1·cm-2	Calibration data, / min-1·cm-2
$21-90	$21-00
$21-91	$21-01
$21-92	$21-02
$21-93	$21-03
$21-94	$21-04

BDPS-02 (γ)
Dose rate at check point, / {dot {H}} rsub {0 i} (10 )	Calibration data, / {bar {dot {H}}} rsub {i} (10)
0,7 µSv/h	$22-00 µSv/h
7 µSv/h	$22-01 µSv/h
70 µSv/h	$22-02 µSv/h
0,7 mSv/h	$22-03 mSv/h
7,0 mSv/h	$22-04 mSv/h
20 mSv/h	$22-05 mSv/h

$eng_certificate_text$.

Calibrated by:	$calibratedBy
		(signature)
Technical control:	$controlledBy
		(signature)